| Physics and Astronomy Dept. | BNB |
| --- | --- |
| UBC | 03/12/09 |
| Multi-Channel Electronics | SC2_ELE_S580_515 |

Multi-Channel Electronics Command Descriptions

# Revision History:

| Rev | Date | Description of change |
| --- | --- | --- |
| 1.0 | 2007-10-09 | MA/added the table to 09/10/07 version |
| 1.1 | 2007-10-18 | BB. Added placeholders for newly supported commands. |
| 1.2 | 2008-01-30 | added scratch and support for slot_id and card_type |
| 1.3 | 2008-05-16 | MA: Updated readout-card command descriptions, references, summary, section 4. Updated command format to not refer to DAS/MAS options as they may change. Added firmware revision information. |
| 1.4 | 2008-05-20 | BB: Added content that I had been working on in my local copy, in CVS. Also BB: Re-integrated information that had been deleted in the update from 1.2 to 1.3. This information has now been reviewed, and is being edited as agreed upon by BB/MA. BB: Added firmware revision columns for all commands BB: Added const_mode and const_data commands for address card. BB: Updated the ‘Command Format’ section so that it is current MA: added servo_mode per column, resolution for gainpid, details for fw_rev, data_mode and captr_raw commands. MA: merged rcs commands with single-rc commands |
| 1.5 | 20080630 | MA added row_dly minimum=4 Added a note about st and go command only supported for ret_dat so far. |
| 1.6 | 20091005 | MA/BB added following commands: num_rows_reported, num_cols_reported, cards_present, cards_to_report, rcs_to_report_data, stop_dly, bias_start, data_mode 10 to 12, readout_col_index, readout_priority. |
| 1.7 | 20091113 | BB: Added the i_clamp_val command and added a note to the const_mode command. Also added run_id and user_writable commands to the clock card section. |
| 1.8 | 20100112 | BB: Amended i_clamp_val, and added mls_sequence_len, mls_data, mls_addr. |

# Summary

This document provides descriptions for commands that can be issued to the Multi-channel Electronics (MCE).  A pair of fibre optic cables connects the MCE to a PC that has a PCI ARC-64 card installed. The PC runs the data acquisition software called DAS (developed at ATC for SCUBA-2 project) or MAS (developed at UBC for non-SCUBA-2 projects using MCE). The user would use MAS or DAS environment to issue commands to the MCE. In DAS, the MCE commands and their encodings are listed in a file called mce.xml while in MAS, such information are in mce.cfg (as of 2008-05-16).

# References

[1] SCUBA2 Bus Backplane ISA, SC2_ELE_S580_511

[2] MCE Firmware Block Diagrams:  SC2/ELE/S563/100, SC2/ELE/S563/200, SC2/ELE/S563/201, SC2/ELE/S563/300, SC2/ELE/S563/400, SC2/ELE/S563/500.

[3] SCUBA2 Data Acquisition Software Overview, Xiaofeng Gao, sc2-sof-s200-014-v1.pdf

[4] SCUBA-2 Data Acquisition to Data Processing Interface, Dennis Kelly, sc2-sof-s200-008rev.pdf

[5] Address-Card Technical Description, SC2_ELE_S584_501

[6] Bias-Card Technical Description, SC2_ELE_S583_501

[7] Clock-Card Technical Description, SC2_ELE_S581_501

[8] Readout-Card Technical Description, SC2_ELE_S582_501

# Command Format

Commands are issued to the MCE from a PC with data acquisition software installed.  At the time of writing, two types of data acquisition software exist:  DAS and MAS.  For both, the nominal command format is the same, but each requires slightly different setup routines before a user can send commands to the MCE.

## DAS

To initiate DAS for commanding the MCE, a user must run two setup scripts:

> cd das

> source dassetup

> das_init start

Following these, a user can issue commands the MCE using the following command format:

> mcecmd <logfile> <action> <card_address> <parameter_id> <values>

i.e. > mcecmd logfile wb cc led 7

## MAS

To initiate MAS for commanding, a user must issue the following command:

> mce_cmd -i

Following this, a user can issue commands the MCE using the following command format:

> <action> <card_address> <parameter_id> <values>

i.e. > wb cc led 7

## The <action> Field

In the MCE command formats for DAS and MAS shown above, the <action> field can take the following values:

- wb: write block – used to write values to registers in the MCE

- rb:  read block – used to read values from registers in the MCE

- go:  go – used to start a process in the MCE that requires several wishbone transactions (i.e. takes and extended period of time)

- rs:  reset – used to issue commands for which replies must be returned by the MCE before executing the original command

- st: stop – used to stop a process that was started with the GO moniker.  Note that “st” is implemented for a ret_dat commands as of 20080630.

## The <card_address> Field

In the MCE command formats for DAS and MAS shown above, the <card_address> field can take the following values.  Each command targets either a particular card or a group of cards:

Single Cards:

- psc: Power Supply Card

- cc:  Clock Card

- rc1:  Readout Card 1

- rc2:  Readout Card 2

- rc3:  Readout Card 3

- rc4:  Readout Card 4

- bc1:  Bias Card 1

- bc2:  Bias Card 2

- bc3:  Bias Card 3

- ac:  Address Card

Card Groups:

- rcs:  All Readout Cards

- bcs:  All Bias Cards

- sys:  All Cards

## The <parameter_id> Field

In the MCE command formats for DAS and MAS shown above, the <parameter_id> field can take the values that are outlined in the section below.

# Command Descriptions

The following table is a summary of commands that can be issued to the MCE.  The commands are divided into sections, depending on what cards may be addressed.  Note that RS- and GO-type commands have special reply-data – which are not listed in the “Reply Data” columns below.  Note also that each of the card types (i.e. Address/ Bias/ Readout/ Clock Cards) in the MCE run unique firmware. The “firmware revision” column indicates the version the particular command is supported.

## General Card Commands

These commands can be issued to any card in the MCE (cc, rc1 to rc4, bc1, bc2, bc3, ac), except the Power Supply Card.  Dark-shaded rows indicate commands that remain to be implemented as of 16 May 2008.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| fpga_temp | 0x91 | rb | N/A | <fpga temperature> | Read the FPGA temperature of a card. | All |
| card_temp | 0x92 | rb | N/A | <card temperature> | Read the card temperature from a IC sensor | All |
| card_id | 0x93 | rb | N/A | <card id> | Read Card ID | All |
| card_type | 0x94 | rb | N/A | <card type> | Request Card Type: Card Type 0: AC Card Type 1: BC Card Type 2: RC Card Type 3: CC Card Type 4: PSUC | All |
| slot_id | 0x95 | rb | N/A | <slot id> | Request Slot ID. | All |
| fw_rev | 0x96 | rb | N/A | <RRrrBBBB> | Read firmware revision where the format in hex is RRrrBBBB RR: major revision number rr: minor revision number BBBB: build number | All |
| dip | 0x97 | rb | N/A | <dip-switch status> | Request DIP-switch status |  |
| led | 0x99 | wb, rb | <status> | <status> | Read/write the LEDs status. When writing to the LEDs, the value written is XOR’ed with the current value of the LED status, to produce the new status. | All |
| scratch | 0x9A | wb, rb | <word0> to <word7> | <word0> to <word7> | These are read/write registers to be used for arbitrarily purposes. An example is to use one to detect undesired reset of MCE. | All |
| vfy_eeprom | 0x42 | wb, rb | <1 = verify> |  | Trigger a comprehensive test of the EEPROM on any/all cards it is addressed to |  |
| clr_error | 0x43 | wb, rb | <1 = reset> |  | Clear error codes that have been recorded on a card. |  |
| resync | N/A | wb, rb | <1 = reset> |  | Tell the targeted card to re-synchronize the start of its internal frame structure with the reception of the next Sync pulse. |  |
| eeprom | 0x41 | wb, rb | <word 0> <word 1> … <word n-1> | <word 0> <word 1> … <word n-1> | Read/write ‘n’ 32-bit words to the EEPROMs on any/all cards starting from the specified starting address (see EEPROM Start). |  |
| eeprom_srt | 0x44 | wb, rb | <address> | <address> | Specify an address in EEPROM memory where a subsequent instruction will write to or read from. The EEPROMs are AT25128A and are 128 KB in size. A 128 KB device holds 4096 32-bit words, so physical addresses for the EEPROM will range between 0x0000 and 0x1000. |  |

## “sys” Commands

These commands affect the timing of the system – which means that they must be issued to all cards in the MCE at the same time.  They are addressed to every card in the card cage by using the ‘sys’ card address.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| row_len | 0x30 | wb, rb | <# clock cycles> | <# clock cycles> | Specify the number of 50MHz clock cycles that are spent per row during multiplexing. Default = 64. | All |
| num_rows | 0x31 | wb, rb | <# rows> | <# rows> | Specify the number of rows to be multiplexed (that there in a detector array). Default = 41. | All |
| num_rows_reported | 0x55 | wb, rb | <number of rows> | <number of rows> | Specify the number of rows of data that are reported in a data packet. This parameter is set in conjunction with the readout_row_index (for Readout Cards). See also num_cols_reported, readout_col_index. Default = 41. | 4.0.0+ |
| num_cols_reported | 0xAD | wb | <# of columns> | <# of columns> | Specify the number of columns of data that are reported in a data packet. This parameter is set in conjunction with the readout_col_index (a Readout Card command). See also num_rows_reported, readout_row_index. Default = 8. | 5.0.0+ |

Figure 2.  Frame Timing Structure (Reference:  \scuba2_repository\doc\ISA\frame_timing_structure.vsd, CVS version 1.6).

Figure 2, above, shows the frame-timing structure of the MCE.  All the commands that affect timing in the MCE are listed in red.  ‘row_len’ and ‘num_rows’ are addressable to all cards in the MCE using the ‘sys’ moniker.  ‘sample_dly’, ‘sample_num’, ‘fb_dly’ are addressable to ‘rcs,’ and ‘row_dly’ is addressable to ‘ac’.

## “psc” Commands

These commands are for the Power Supply Card only.  The Power Supply Card firmware blocks are integrated in the Clock Card firmware as a virtual card.  This is why the firmware revision refers to the Clock Card.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| brst | 0x60 | rs | 1 | N/A | Cause the Power Supply Card to assert the RBst line (active low) on the bus backplane to reset all FPGA devices in the MCE. Asserting this signal causes all FPGA devices in the MCE to be reconfigured, including the Clock Card FPGA. The Clock Card FPGA will be reconfigured by the Factory reconfiguration device. | 2.0.c+ |
| cycle_pow | 0x61 | rs | 1 | N/A | Cause the Power Supply Card to cycle power to the MCE. All power supplies will be switched off and then back on again in proper sequence. This will cause all FPGA devices in the MCE to be reconfigured, including the Clock Card FPGA. The Clock Card FPGA will be reconfigured by the Factory configuration device. | 2.0.c+ |
| psc_status | 0x63 | rb | N/A | <status word 0> <status word 1> … <status word 35> | Report power-supply-card status information. | 2.0.c+ |
| cut_pow | 0x62 | rs | 1 | N/A | Cause the Power Supply Card to turn off all power bricks. | 2.0.c+ |

## “cc” Commands

These commands are for the Clock Card only.  Dark-shaded rows indicate commands that remain to be implemented as of 16 May 2008.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| config_app | 0x52 | rs | 1 | N/A | Cause the Clock Card’s application-configuration device to load its configuration data into the Clock Card FPGA. | 3.0.1+ |
| config_fac | 0x51 | rs | 1 | N/A | Cause the Clock Card’s factory-configuration device to load its configuration data into the Clock Card FPGA. | 3.0.1+ |
| ret_dat_s | 0x53 | wb, rb | <start sequence #> <end sequence #> | <start sequence #> <end sequence #> | Specify the starting and ending sequence-numbers for data frames that the MCE is to return. Each frame of data will have a sequence number affixed to it. Sequence numbers will be assigned to data frames in increasing order, beginning with the starting sequence number, and finishing with the ending sequence number. Both parameters may range between 0x00000000 and 0xFFFFFFFF. If both starting and ending sequence numbers are the same, then a single frame of data will be returned. If the starting sequence number is larger than the ending sequence number, then frames will be returned until the sequence number has wrapped around and reached the stopping sequence number. Starting and stopping sequence numbers are expected for all modes of data acquisition – including for DV pulses. | All |
| use_dv | 0x54 | wb, rb | <dv mode> | <dv mode> | Set the dv mode. The Clock Card can return data frames at an internal rate specified by ‘data_rate’, or in response to Data-Valid pulses from the fibre input or in response to packets from the Manchester input. Default = 0. dv mode 0: generated by the MCE dv mode 1: fibre DV input (labeled as DV_IN on the MCE) dv mode 2: fibre Manchester-encoded input (labeled as SYNC_IN on the MCE) use_sync, use_dv and select_clk must be consistent with each other, otherwise timing artifacts will appear in data. To be consistent, they must either be [0,0,0] or [2,2,1] respectively. | All |
| array_id | 0x58 | rb |  | <array id> | Read the identification of the sub-array that a MCE box is positioned for connection to, on the side of the Dewar. | All |
| box_id | 0x59 | rb |  | <id word 1> | Read the identification of the sub-rack. Each production sub-rack will have a unique Silicon ID IC mounted on the Bus Backplane. | All |
| sram_data | 0x5C | wb, rb | <word 0> … <word n> | <word 0> … <word n> | Read/write a block of 32b data to SRAM at address specified by sram_addr parameter.(see sram_addr ) | 4.0.9+ |
| sram_addr | 0x5E | wb, rb | <starting address> | < starting address> | Specify the starting address for subsequent read/write operation to the 1024k x 32-bit SRAM. The, physical addresses range between 0x00000000 and 0x000FFFFF. | 4.0.9+ |
| data_rate | 0xA0 | wb, rb | <# frame periods per data frame> | <# frame periods per data frame> | Specify the rate at which data frames are collected by the MCE if DV pulses are not being used (see “Use DV Pulse”). The rate is specified as the time between data packets measured in frame periods, i.e. a value of 11 will return one data packet every 11 frame periods. A frame period is the amount of time required for the multiplexer to address all the rows on a MUX. Default = 47. | All |
| use_sync | 0xA1 | wb, rb | <sync mode> | <sync mode> | Specify the sync mode. The Clock Card can return to address 0 using its internal timing, or in response to sync pulses from the fibre input or in response to packets from the Manchester input. Default = 0. sync mode = 0: generated by the MCE sync mode = 1: fibre DV input (labeled as DV_IN on the MCE) sync mode = 2: fibre Manchester-encoded input (labeled as SYNC_IN on the MCE) use_sync, use_dv and select_clk must be consistent with each other, otherwise timing artifacts will appear in data. To be consistent, they must either be [0,0,0] or [2,2,1] respectively. | All |
| select_clk | 0xA2 | wb, rb | <clock source> | <clock source> | Select the source of the clock input to the Clock Card’s PLL. There are two options for this, either the Manchester input clock, or the on-board crystal clock. Both are nominally 25 MHz. Default = 0. clock source = 0: generated by the MCE clock source = 1: fibre Manchester-encoded input. (labeled as SYNC_IN on the MCE) Note: when select_clk=1, but the external Manchester Clock disappears, the firmware automatically falls back on the internal clock and reading back select_clk would reflect that as 0. use_sync, use_dv and select_clk must be consistent with each other, otherwise timing artifacts will appear in data. To be consistent, the must either be [0,0,0] or [2,2,1] respectively. | All |
| upload_fw | 0x50 | wb, rb | <word 0> <word 1> … <word 58> | <word 0> <word 1> … <word 58> | Write data to the JAM player wishbone slave, which will stream the words over the JTAG chain to program a device. | Not yet implemented |
| mce_bclr | 0xAB | rs | 1 | N/A | Clear the registers on all the cards in the MCE, including the Clock Card but excluding the PSC. | 3.0.4+ |
| cc_bclr | 0xAC | rs | 1 | N/A | Clear the registers on the Clock Card. | 3.0.4+ |
| tes_tgl_en | 0xA3 | wb, rb | 1 | <mode> | Enable or disable a square wave applied to the TES Bias. When this mode is enabled, the Clock Card issues internal commands at timed intervals that toggle the TES Bias output between a low value and a high value. mode = 0: Disable TES Bias toggling mode = 1: Enable TES Bias toggling | Until 3.0.8 |
| tes_tgl_max | 0xA4 | wb, rb | 1 | <high value> | Specify the ‘high’ value of the TES Bias square wave | Until 3.0.8 |
| tes_tgl_min | 0xA5 | wb, rb | 1 | <low value> | Specify the ‘low’ value of the TES Bias square wave | Until 3.0.8 |
| tes_tgl_rate | 0xA6 | wb, rb | 1 | <toggle rate> | Specify the half-period of the TES Bias square wave, in frame periods | Until 3.0.8 |
| int_cmd_en | 0xA7 | wb, rb | 1 | <mode> | Enable or disable internal housekeeping commands. Default = 0. mode = 0: Disable internal commands mode = 1: Enable internal commands | Until 4.0.0 |
| box_temp | 0xA8 | rb | 1 | <temperature> | Read the temperature sensor on the Bus Backplane | All |
| crc_err_en | 0xA9 | wb, rb | 1 | <mode> | Force the value of the CRC field in reply packets from the MCE to the PC. The reply packets will have a CRC checksum value of 0xFFFFFFFF. mode = 0: Normal CRC Calculation performed mode = 1: CRC word forced to 0xFFFFFFFF | 3.0.1+ |
| internal_cmd_mode | 0xB0 | wb, rb | 1 | <mode> | Enable or disable internal housekeeping/ ramping commands. Default = 0. mode = 0: Disable internal commands mode = 1: Enable Housekeeping commands mode = 2: Enable ramp commands mode = 3: Enable maximum length sequence (MLS) commands | 4.0.0+ |
| ramp_step_period | 0xB1 | wb, rb | <# frame periods> | <# frame periods> | Specify the number of frame periods between each ramp step. | 4.0.0+ |
| ramp_min_val | 0xB2 | wb, rb | <min value> | <min value> | Specify the minimum value of the ramp that is to be applied. Must be less than the ramp_max_value | 4.0.0+ |
| ramp_step_size | 0xB3 | wb, rb | <step size in DAC units> | <step size in DAC units> | Specify the ramp step size. If (step_size) = [(max) – (min)] then the ramp will be a square wave. | 4.0.0+ |
| ramp_max_val | 0xB4 | wb, rb | <max value> | <max value> | Specify the maximum value of the ramp that is to be applied. Must be greater than the ramp_min_value | 4.0.0+ |
| ramp_param_id | 0xB5 | wb, rb | <parameter id> | <parameter id> | Specify the parameter ID of the register to be ramped | 4.0.0+ |
| ramp_card_addr | 0xB6 | wb, rb | <card address> | <card address> | Specify the card address of the register to be ramped | 4.0.0+ |
| ramp_step_data_num | 0xB7 | wb, rb | <number of values to ramp> | <number of values to ramp> | Specify the number of data that are to be written per ramp command | 4.0.0+ |
| config_jtag | 0xAA | rs | 1 | N/A | Trigger the configuration of a specific configuration-device over JTAG using a file stored in the SRAM on the Clock Card. The “Configure MCE Device Setup” command must be issued before this one to specify which device the Clock Card’s programmer is to target with the configuration file. Possible devices include the Clock Card’s application configuration device, as well as configuration devices on any other cards (except the Power Supply Card, which has none). |  |
| cards_present | 0x5A | rb | 1 | <cards present> | Read which cards are present in the MCE. Bits are one-hot encoded, active high: Bit 9: Address Card Bit 8: Bias Card 1 Bit 7: Bias Card 2 Bit 6: Bias Card 3 Bit 5: Readout Card 1 Bit 4: Readout Card 2 Bit 3: Readout Card 3 Bit 2: Readout Card 4 Bit 1: Clock Card Bit 0: PSUC | 4.0.8+ |
| cards_to_report | 0x5B | wb, rb | <cards to report data> | <cards to report data> | Specify which cards are present in the MCE subrack, and will return data over the bus backplane (whether real or dummy). This command can be used in situations where certain cards are not present in the subrack, and should not return data in replies. The default value for this register is ‘1’ for every possible card. To stop a card from returning data, set its corresponding bit to ‘0’. Bit 9: Address Card Bit 8: Bias Card 1 Bit 7: Bias Card 2 Bit 6: Bias Card 3 Bit 5: Readout Card 1 Bit 4: Readout Card 2 Bit 3: Readout Card 3 Bit 2: Readout Card 4 Bit 1: Clock Card Bit 0: PSUC | 4.0.a+ |
| ret_dat_req | 0x5D | wb, rb | <1 = enable; 0 = disable> | <1 = enable; 0 = disable> | The user can start or stop the data acquisition process on the Clock Card. Not supported yet. |  |
| rcs_to_report_data | 0x5F | wb, rb | <card address> | <card address> | This command is NOT redundant with cards_to_report. It applies only to Readout Cards, and only for data taking (ret_dat) commands. This register sets which Readout Cards return data during a data run. This command allows the selective return of data without affecting which cards respond to all other commands, which is useful if a data run fails. Bit 9: -- Bit 8: -- Bit 7: -- Bit 6: -- Bit 5: Readout Card 1 Bit 4: Readout Card 2 Bit 3: Readout Card 3 Bit 2: Readout Card 4 Bit 1: -- Bit 0: -- | 4.0.a+ |
| stop_dly | 0xB8 | wb, rb | <timed delay> | <timed delay> | Specify the time delay between the return of the next data frame and the return of the reply to a stop command. The time specified is in microseconds. This command is used to test the robustness of the PCI card to replies following immediately on the heels of a data packet. | 4.0.a+ |
| run_id | 0x56 | wb, rb | <data file name embedded in header> | <data file name embedded in header> | Specify an ID number that will be stored in every data packet header returned to the PCs. In ACT, the ID number corresponds to the c-time at which data acquisition began. ACT’s data files are also named by the same c-times, so that the data pipeline can easily track which files the data are from. | 4.0.2+ |
| user_writable | 0x57 | wb, rb | <misc data embedded in header> | <misc data embedded in header> | Specify a value that will be stored in every data packet header returned to the PCs. This register is used by ACT to store a combination of array_id and data_mode information. At Caltech, it is used during I-V curves to store the value of the TES bias applied. | 4.0.2+ |
| mls_sequence_len | 0xB9 | wb, rb | <sequence length> | <sequence length> | Specify the length of the MLS sequence that is loaded in RAM. This parameter tells the internal commanding FSM how long the MLS sequence is, and at what index it should restart MLS the sequence. | 5.0.3+ |
| mls_data | 0xBA | wb, rb | <data word 0> <data word 1> … <data word n> | <data word 0> <data word 1> … <data word n> | Load the MLS sequence into RAM on the Clock Card. The RAM has 8192 indexes. This may be expanded later if there are enough resources still available on the Clock Card. | 5.0.3+ |
| mls_addr | 0xBC | wb, rb | 1 | 1 | Specify the starting address for reading from or writing to the MLS RAM. This parameter does not affect the memory address of the internal commanding FSM. | 5.0.3+ |

## “ac” Commands

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| row_order | 0x01 | wb, rb | <1st row> <2nd row> … <41st row> | <1st row> <2nd row> … <41st row> | This command is relevant when enbl_mux = 1 or 2. Read/write the row-addressing order. A sequence of up to 41 rows is specified at once. The row numbers that are specified with this command refer to the physical channel numbers on the Address Card. Note: After this command is issued, a “flx_lp_init” command must be issued to all the readout cards to discard previous PID-loop calculations and to clean out the rest of the data pipeline. Default = 0,0,0,… | All |
| on_bias | 0x02 | wb, rb | <bias row 0> <bias row 1> … <bias row 39> <bias dark row> | <bias row 0> <bias row 1> … <bias row 39> <bias dark row> | This command is relevant when enbl_mux = 1. Read/write the14-bit on-bias values for all 41 channels of the DACs on the Address Card. Bias values are stored in the order of physical channels on the Address card – which is not necessarily the multiplexing order specified with row_order. This command is nominally is used to specify the 1st Stage SQUID “on” biases, and are applied one at a time in the sequence specified by row_order. Default = 0,0,0,… | All |
| off_bias | 0x03 | wb, rb | <bias row 0> <bias row 1> … <bias row 39> <bias dark row> | <bias row 0> <bias row 1> … <bias row 39> <bias dark row> | This command is relevant when enbl_mux = 1. Read/write the 14-bit off-bias values for all 41 channels of the row-addressing DACs on the Address Card. Bias values are stored in the order of physical channels on the Address card – which is not necessarily the multiplexing order specified with row_order. This command is nominally is used to specify the 1st Stage SQUID “off” biases. Default = 0,0,0,… | All |
| enbl_mux | 0x05 | wb, rb | <mode> | <mode> | Start/stop array multiplexing. Default: mode = 0. mode = 0: Multiplexing off. This is the default value. mode = 1: (changed starting in v2.0.5+) Multiplexing on. The Address Card turns one DAC ‘on’ and one DAC ‘off’ at every row switch. The multiplexed values are specified by on_bias and off_bias. In v2.0.4-, the values were applied starting at row_dly, and done by row_dly+4. Starting in v2.0.5+, the values are applied in the first two clock cycles after a row switch, and row_dly is obsolete. mode = 2: Multiplexing on. The Address Card changes the values on all the DACs at every row switch. The multiplexed values are specified by fb_colx. This mode is implemented starting in v2.0.5+. All 41 DAC values are applied in the first 4 clock cycles after a row switch, and row_dly is obsolete. | All |
| bias_start | 0x09 | wb, rb | <start row 0> <start row 1> … <start row 39> <start dark row> | <start row 0> <start row 1> … <start row 39> <start dark row> | Applies only when enbl_mux = 1. Specify the point during each row dwell period when the 1st Stage Bias is applied. This is specified on a row-by-row basis so that SCUBA2 can differential bias heating across the rows of their arrays. | 2.0.8; 5.0.1+ |
| row_dly | 0x35 | wb, rb | <# clock cycles> | <# clock cycles> | This command is obsolete. This command was in use until v2.0.4 for enbl_mux = 1 only. Specify the number of 50MHz clock cycles between the de-assertion of the previous row at a row-switch, and the start of the assertion of the current row on the Address Card (multiplexer). Minimum row_dly is 4. This command is not supported after 2.0.5 version. | 2.0.4-; |
| fb_col0 … fb_col40 | 0xC0 … 0xE8 | wb, rb | <fb_row_0> <fb_row_1> … <fb_row_39> <fb_row_dark> | <fb_row_0> <fb_row_1> … <fb_row_39> <fb_row_dark> | This command is relevant when enbl_mux = 2. Specifies the 41 multiplexing values for each DAC channel on the Address Card, and for each row that the channel is to multiplex. There are 41 channels on the Address Card, and up to 41 rows per channel that need to be specified. This command was implemented to support using the Address Card to multiplex the 2nd Stage Feedback from a Bias Card slot. | 2.0.5+ |
| const_mode | 0x06 | wb, rb | <mode_dac_0> <mode_dac_1> … <mode_dac_39> <mode_dac_dark> | <mode_dac_0> <mode_dac_1> … <mode_dac_39> <mode_dac_dark> | Specify which DAC channels are held at a constant value regardless of the value of enbl_mux. A constant value is applied once only, and is specified by const_val. In addition, DACs for which const_mode=1 do not receive data strobes during multiplexing – so that there are no glitches or excess noise on the DAC outputs. mode_dac_x = 0: this DAC channel is multiplexed when enbl_mux = 1, or enbl_mux = 2, or enbl_mux=3 mode_dac_x = 1: the corresponding DAC channel is held at a constant value and not strobed when enbl_mux = 1, or enbl_mux = 2. | 2.0.6+ |
| const_val | 0x07 | wb, rb | <val_dac_0> <val_dac_1> … <val_dac_39> <val_dac_dark> | <val_dac_0> <val_dac_1> … <val_dac_39> <val_dac_dark> | Specify values for the DACs that are held constant. <val_dac_n> is asserted by DAC ‘n’ when const_mode[n] = 1. In addition, all constant values are applied to each respective DAC when enbl_mux = 0. When <wb enbl_mux>, <wb const_mode>, <wb const_val> or <wb const_val39> commands are executed, all the DACs that are being held constant have their values re-strobed. This command was implemented to avoid strobing DACs that are connected to dead rows or columns, while multiplexing the rest. This effectively eliminates the noise in the cryostat associated with strobing DACs. | 2.0.6+ |
| const_val39 | 0x08 | wb, rb | <val_dac_39> | <val_dac_39> | Specify the value that DAC 39 is to be held constant at. This value will be asserted if const_mode[39] = 1, or if enbl_mux = 0. This command was created to allow internal commands to ramp this DAC directly, and to avoid excess strobe noise, because it is connected to the TES Bias in some ACT cryostats. | 2.0.7+ |

## “rc1, rc2, rc3, rc4” Commands

These commands target the Readout Cards only.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware Revision |
| --- | --- | --- | --- | --- | --- | --- |
| sa_bias | 0x10 | wb, rb | <column 0 SA bias> <column 1 SA bias> … <column 7 SA bias> | <column 0 SA bias> <column 1 SA bias> … <column 7 SA bias> | Read/write the SQUID Series Array Bias values for all columns on a Readout Card. | All |
| offset | 0x11 | wb, rb | <column 0 offset> <column 1 offset> … <column 7 offset> | <column 0 offset> <column 1 offset> … <column 7 offset> | Read/write the SQUID Series Array Offset values for all columns on a Readout Card. | All |
| gainp0 gainp1 gainp2 gainp3 gainp4 gainp5 gainp6 gainp7 | 0x70 0x71 0x72 0x73 0x74 0x75 0x76 0x77 | wb, rb* | <column x, gainp 0> <column x, gainp 1> … <column x, gainp 38> <column x, gainp 39> <column x, gainp dark> | <column x, gainp 0> <column x, gainp 1> … <column x, gainp 38> <column x, gainp 39> <column x, gainp dark> | Read/write the Readout Card’s 41 P coefficients for any given column. The <x> in the command moniker refers to which column the values are being specified for, from 0 to 7. The Parameter IDs are listed in order of column, from 0 to 7. gainp are specified as “signed 10-bit” values as of revision 4.0.2. (These were 8-bit values prior to revision 4.0.2.) | All * |
| gaini0 gaini1 gaini2 gaini3 gaini4 gaini5 gaini6 gaini7 | 0x78 0x79 0x7A 0x7B 0x7C 0x7D 0x7E 0x7F | wb, rb* | <column x, gaini 0> <column x, gaini 1> … <column x, gaini 38> <column x, gaini 39> <column x, gaini dark> | <column x, gaini 0> <column x, gaini 1> … <column x, gaini 38> <column x, gaini 39> <column x, gaini dark> | Read/write the Readout Card’s 41 I coefficients for any given column. The <x> in the command moniker refers to which column the values are being specified for, from 0 to 7. The Parameter IDs are listed in order of column, from 0 to 7. gaini are specified as “signed 10-bit” values as of revision 4.0.2. (These were 8-bit values prior to revision 4.0.2.) | All * |
| flx_quanta0 flx_quanta1 flx_quanta2 flx_quanta3 flx_quanta4 flx_quanta5 flx_quanta6 flx_quanta7 | 0x80 0x81 0x82 0x83 0x84 0x85 0x86 0x87 | wb, rb | <column x, quanta 0> <column x, quanta 1> … <column x, quanta 38> <column x, quanta 39> <column x, quanta drk> | <column x, quanta 0> <column x, quanta 1> … <column x, quanta 38> <column x, quanta 39> <column x, quanta drk> | Read/write the Readout Card’s 41 Flux Quanta values for any given column. The Flux Quanta are used to calculate the 1st Stage Feedback value. The <x> in the command moniker refers to which column the values are being specified for, from 0 to 7. The Parameter IDs are listed in order of column, from 0 to 7. | All |
| gaind0 gaind1 gaind2 gaind3 gaind4 gaind5 gaind6 gaind7 | 0x88 0x89 0x8A 0x8B 0x8C 0x8D 0x8E 0x8F | wb, rb* | <column x, gaind 0> <column x, gaind 1> … <column x, gaind 38> <column x, gaind 39> <column x, gaind dark> | <column x, gaind 0> <column x, gaind 1> … <column x, gaind 38> <column x, gaind 39> <column x, gaind dark> | Read/write the Readout Card’s 41 D coefficients for any given column. The <x> in the command moniker refers to which column the values are being specified for, from 0 to 7. The Parameter IDs are listed in order of column, from 0 to 7. gaind are specified as “signed 10-bit” values as of revision 4.0.2. (These were 8-bit values prior to revision 4.0.2.) | All * |
| adc_offset0 adc_offset1 adc_offset2 adc_offset3 adc_offset4 adc_offset5 adc_offset6 adc_offset7 | 0x68 0x69 0x6A 0x6B 0x6C 0x6D 0x6E 0x6F | wb, rb | <column x, offset 0> <column x, offset 1> … <column x, offset 38> <column x, offset 39> <column x, offset dark> | <column x, offset 0> <column x, offset 1> … <column x, offset 38> <column x, offset 39> <column x, offset dark> | Read/write the Readout Card’s 41 ADC Offsets for any given column. The ADC Offset coefficients are used to digitally offset each 50 MHz ADC sample on the Readout Cards. The offset is subtracted from each 50 MHz ADC sample before being co-added. The <x> in the command moniker refers to which column the values are being specified for, from 0 to 7. The Parameter IDs are listed in order of column, from 0 to 7. | All |
| readout_row_index | 0x13 | wb, rb | <starting row index> | <starting row index> | Specify the starting row index of data to be returned in a data block. Default = 0. Note: if num_rows_reported > num_rows – readout_row_index – 1, then the readout row index will wrap accordingly to 0 during readout. | Starting 4.0.0 |
| ret_dat | 0x16 | go, st | <1 = return frames> |  | Request data frames. Prior to a ret_dat command, the Clock Card expects a “ret_dat_s” command to set up how many frames are to be collected and what sequence numbers to assign. Following that, a single “ret_dat” command triggers the return of n frames of data. | All |
| fltr_rst | 0x14 | wb, rb | <1 = reset filter pipeline> | <1 = reset filter pipeline> | Reset the filter data pipeline on the Readout Cards. The flx_lp_init command automatically triggers an fltr_rst. | All |
| en_fb_jump | 0x15 | wb, rb | <1 = enable flux jumping> | <1 = enable flux jumping> | Enable/disable flux jumping on the Readout Cards. The size of flux jumps is specified by the ‘flux_quanta’ command. | All |
| data_mode | 0x17 | wb, rb | <data mode> | <data mode> | Specify the pixel-data read out mode. This determines whether the coadded-error, the first-stage feedback, the filtered feedback, the flux-jump counter, or some combination of the above would be read. (Consult readout card description document [8] for updated info and windowing information of each mode.) Some of the available modes are: Data mode 0: Signed 32-bit co-added error data: error[31:0] Data mode 1: signed 32-bit first-stage feedback data (fb) , when servo_mode=3 this is PID calculation result and the actual value applied to the DAC is fbDAC=fb/212: fb[31:0] Data mode 2: signed 32-bit low-pass filtered first-stage feedback data: fltr[31:0] Data mode 3: raw data sampled at 50MHz (see captr_raw command): raw[13:0] Data mode 4: mixed data: 18b SQ1_fb + 14b coadded error: fb[31] & fb[28:12] & coadd[31] & coadd[12:0] Data mode 5: mixed data: 24b SQ1_fb + 8b flux-jump counter: fb[31:8] & fj[7:0] Data mode 6: mixed data: 18b filtered fb + 14b coadded error: fltr[31] & fltr[27:11] & error[31] & error[12:0] Data mode 7: mixed data: 22b filtered fb + 10b coadded error: fltr[31] & fltr[27:7] & error[31] & error[12:4] Data mode 8: mixed data: 24b filtered fb + 8b flux-jump counter: fltr[31:8] & fj[7:0] Data mode 9: mixed data: 24b filtered fb + 8b flux-jump counter: fltr[31] & fltr[23:1] & fj[7:0] Data mode 10: mixed data: 25b filtered fb + 7b flux-jump counter: fltr[27:3] & fj[6:0] Data mode 11: pixel addresses: 6 bits of row index + 3 bits of column index: row[8:3] & col[2:0] | See Description Pane (Left) 0: All 1: All 2: 1.4.1+ 3: Special Req. 4: All 5: All 6: 3.0.30 – 4.0.6 7: 4.0.2+ 8: 4.0.4 only 9: 4.0.5 – 4.0.a 10: 4.0.b+ 11: 5.0.0+ |
| captr_raw | 0x18 | wb, rb | <1 = capture data> | N/A | Initiates taking a snap-shot of 50MHz ADC samples for 2 consecutive frames and storing them in raw-mode buffer on the Readout Card. This command is only available in special builds of firmware. (In 4.3.7, the raw-buffer is 14-bit wide which amounts to storing full-resolution 50MHz ADC samples.) | 3.0.19, 4.1.7, 4.2.7, 4.3.7 Only |
| servo_mode | 0x1B | wb, rb | <column 0> ….. <column 7> | <column 0> …. <column 7> | Set the servo on the Readout Cards. Servo modes include: Servo mode 0 or 1: constant feedback mode as specified by fb_const parameter Servo mode 2: ramp mode Servo mode 3: PID-loop lock mode (servo parameters vary depending on firmware build) | All |
| ramp_dly | 0x1C | wb, rb | <# frames> | <# frames> | Specify the number of frame-periods before the next step in a ramp function. This delay can be tuned to the data-frame readout rate (data_rate), so that one frame of data is read out per ramp step. | All |
| ramp_amp | 0x1D | wb, rb | <ramp amplitude> | <ramp amplitude> | Specify the maximum ramp value, in 14-bits | All |
| ramp_step | 0x1E | wb, rb | <ramp step size> | <ramp step size> | Specify what the step size of each step in the ramp is, max 14-bit | All |
| fb_const | 0x1F | wb, rb | <col0> <col1> <col2> <col3> <col4> <col5> <col6> <col7> | <col0> <col1> <col2> <col3> <col4> <col5> <col6> <col7> | Specify the constant feedback DAC value for a specific column when the servo is in constant mode (see servo_mode). Applied to SQ1_FB. Max =14-bit (-8192 to 8191). The constant value is applied to the Readout Card DAC repetitively for every new row. ( | All |
| sample_dly | 0x32 | wb, rb | <# clock cycles> | <# clock cycles> | Specify the number of 50MHz clock cycles from the start of a row during which samples do not contribute to a co-added value on the Readout Cards | All |
| sample_num | 0x33 | wb, rb | <# samples> | <# samples> | Specify the number of 50MHz samples co-added during a row-period on the Readout Cards. Co-addition begins after the sample delay, and one value is co-added per 50MHz clock cycle. Note that default is 0 and therefore you need to explicitly set this parameter. | All |
| fb_dly | 0x34 | wb, rb | <# clock cycles> | <# clock cycles> | Specify the number of 50MHz clock-cycles between the start of a row and the assertion of the row’s feed-back value on the Readout Cards. (minimum is 7 with flux-jumping off and 18 when flux-jumping is on) | All |
| flx_lp_init | 0x37 | wb, rb | <1 = initialize> | N/A | Initialize/reset the PID loop calculations on the Readout Cards for new PID parameters to take effect. | All |
| filt_coef | 0x1A | wb, rb | <coefficient 1> … <coefficient n> | <coefficient 1> … <coefficient n> | Specify coefficients for low-pass filter. (not implemented) | none |
| readout_col_index | 0x19 | wb, rb | <starting column index> | <starting column index> | Specify the starting column index of data to be returned in a data block. Default = 0. Note: if num_cols_reported > num_cols – readout_col_index – 1, then the readout column index will wrap accordingly to 0 during readout. During raw readout, this parameter specifies the index of the column for which raw data is stored. |  |
| readout_priority | 0x67 | wb, rb | <priority> | <priority> | Specify whether readout should be column-major or row-major. Default = 0. Priority 0: Data readout is in row-major order – all the pixels in a row are read out before moving on to the next row. Priority 1: Data readout is in column-major order – all the pixels in a column are read out before moving on to the next column. | 5.0.0+ |
| i_clamp_val | 0x66 | wb, rb | <clamp_value> | <clamp_value> | Specify the maximum value of the integral term before it is multiplied by the integral coefficient. This register specifies the maximum absolute value of the integral term that is allowed, and clamps the integral term at this value if it is exceeded. This prevents unlocked pixels from wrapping continuously in lock mode and mitigates the adverse effects of this on locked pixels. The value of this register is different for each experiment and depends on the integral coefficient, and flux quanta size. The formula for determining an appropriate value is: i_clamp_val < {[220] * [min_flx_quanta]} / [max_i_coeff] | 5.0.9+ |

*In Pre-4.0.e firmware that has raw-mode enabled read back of gaini, gainp, gaind parameters is disabled to save memory for raw mode.

## “bc1, bc2, bc3” Commands

These commands are for Bias Cards only.  Dark-shaded rows indicate commands that remain to be implemented as of 2008-05-16.

| Parameter ID | Register Address | Possible Actions | Command Parameters | Reply Data | Description | Firmware revision |
| --- | --- | --- | --- | --- | --- | --- |
| flux_fb | 0x20 | wb, rb | <flux fb, column 0> <flux fb, column 1> … <flux fb, column 31> | <flux fb, column 0> <flux fb, column 1> … <flux fb, column 31> | Specify the flux feedback for all 32 16-bit DAC channels on the Bias Cards. In MCEv1, the nominal mappings are: Bias Card 1 ↔ SQUID Series Array Feedback (32 channels), Bias Card 2 ↔ 2nd Stage SQUID Feedback (32 channels), Bias Card 3 ↔ 2nd Stage SQUID Bias (32 channels). In MCEv2, the nominal mappings are: Bias Card 1 ↔ SQUID Series Array Feedback (lower 16 channels), Bias Card 1 ↔ 2nd Stage SQUID Bias (upper 16 channels), Bias Card 2 ↔ 2nd Stage SQUID Feedback (lower 16 channels), Bias Card 3 ↔ TES Bias (lower 16 channels). | All |
| bias | 0x21 | wb, rb | <bias value> | <bias value> | Read/Write the low-noise LVDS 16-bit DACs that are typically used to supply TES bias or pixel heater currents. The nominal mappings are: Bias Card 1 ↔ Pixel Heater, Bias Card 2 ↔ TES Bias. | All |
| flux_fb_upper | 0x24 | Wb,rb | <flux fb, column 16> <flux fb, column 17> … <flux fb, column 31> | <flux fb, column 16> <flux fb, column 17> … <flux fb, column 31> | Specify the flux feedback for 16 upper 16-bit DACs, i.e. channel 16 to 31 on bias cards. This command was implemented to match the cryostat wiring for the MCEv2 subracks, which use the upper 16 channels of BC1 to supply the SQ2 Bias. | 1.4.0 + |
| sa_htr0 | 0x22 | wb, rb | <heater value> | <heater value> | Read/write a from/to the Series Array Heather channel 0. This is used to release trapped flux from the Series Array Modules. | None |
| sa_htr1 | 0x23 | wb, rb | <heater value> | <heater value> | Read/write a heater value from/to the Series Array Heather channel 1. This is used to release trapped flux from the Series Array Modules. | None |